## Slide 1
Kim Feijen
YS375F – Continuous Wavelength Module
Opto-mechanical architect

## Slide 2: Thick stacks make robust recipe determination difficult
Increasing stack thickness increases the swing curve behaviour however the stack complexity can also have an influence on “thin” stacks.
Known memory user cases (eF2X / D2H):
1.1 µm DRAM (A1)
0.9 µm DRAM (C6)
Stack sensitivity
6.5 µm 3D NAND (NONO, 96 tiers)
3 µm DRAM (D5)
Peak-peak: 20 nm  fine WL adjustment required
Sensing wavelength
Recipes using “peak” wavelengths are least sensitive to process variation
Slide MMBM
Notes: Why do we need CL? Our application friends did some measurements with a MC 250 LPPS and next to that simulations. In the results they see that the stack sensitivity is a swing curve which correlates with a wavelength. This swing curve behavior is increasing when the stacks become thicker like we can see on the right picture.

## Slide 3: Current situation to near future
LPPS YS350 → MC LPPS
4x
Notes: If we look to the current situation, then we see that the stack sensitivity for instance on [click] 600nm is zero. This means we have a very bad stack response resulting in difficult to measure markers. [click] At approx. 620nm we see the highest stack sensitivity. For the customer it would be nice if he can select a wavelength which is close to this peak. [click] With the MC LPPS we can give the customer some additional wavelengths already, as we swap the WB filters for new NB filters in the LPPS.

## Slide 4: Tunable wavelength selector
Tunable wavelength to improve stack sensitivity
Notes: Going one step further would be giving the customer all desired wavelengths and band widths between 425 and 900nm. This can be done by making use of linear gradient filters, such are some being supplied by Delta Thin film. [click]

## Slide 5: Tunable wavelength selector
Tunable wavelength to improve stack sensitivity
Linear gradient filters
https://www.youtube.com/watch?feature=player_detailpage&v=BcXJoIOtVKQ#t=12
Physics behind thin film interference
https://youtu.be/xjMjWtntm9k?t=348
Notes: The basic principle is explained in a movie [MOVIE]
The physics principle behind interference filters is explained in the other movie. [MOVIE]

## Slide 6: Design constraints
Commercially available linear variable filters
Restricted volume and mass available in the YS
Fast switching required (< 45ms)
One spare slot in YMCR available
Limited power available
No cooling water available

## Slide 7: Functional breakdown
Modular approach with a balance between functions and manufacturability
Functions broken down in modular design
Modular design balanced on serviceability and manufacturability with process flow

## Slide 8: Baseline lay-out CL module January 10th 2017
MOTOR FRAME
5x ND
1x slot
3x spare
OPTICAL BASE FRAME
LWP
SWP
LPPS
Mixing fiber
Z
Y
≅4%
≅ 4%
Spectrometer
H
V
Cal. source
F1
L1
M1
L2
L3
BS1
M3
L4
F3
F2
F4
L5
BS2
ND1
M2
L6
ND2

## Slide 9: CL Module Process Flow
4022.666.35251
4022.666.35551
4022.666……
4022. 656.13551
4022.666.35451
656. 13381
656. 13371
656. 13271
656. 13261
656. 13281
656. 13391
656. 13421
656. 13431
656. 13241
CL MOD
calibrations
LWP-filter
Bond / assemble
CL MOD
assemble
SWP-filter
Bond / assemble
S
Spare Part Diagnostics
…
Version w1645.2
…
∑8k€
fibers
4022.666.35351
666. 35101
4022.666.35151
4022.656.13231
4022.656.13131
4022.656.13051
4022.666.35201
4022.666.35401
4022.656.13171
4022.656.13181
4022.656.13191
4022.656.13211
4022.656.13141
Shaft.
XWP flt-mtr
assy qualified
CL MOD
qualified
Spectro Branch
qualified
4022.656.13061
LWP-fltr assy
qualified
SWP-fltr assy
qualified
4022.666.35501
CL housing
ND flt-mtr assy qualified
S
R
Reliability Critical Part
R
R
R
S
VQI 1
VQI 2
VQI 3
VQI 3
VQI 3
VQI 3
VQI 3
VQI 3
VQI 3
VQI 3
VQI 2
VQI 0
VQI 0
Spectro meter
Ref.
source
Filter
3 of 4
Filter
4 of 4
Filter
2 of 4
Filter
1 of 4
Wheel frame
R
R
R
R
Diagn: ACPD Sens&Act
Diagn: TBD
Diagn: TBD
Diagn: TBD
Adjust filter to motor & cal
Optical Assy
Adjust & Cal
Optical Assy
assemble
Optical frame
qualified
xxxx
phantom
Highest level buy
Complex Part
xxxx
qualified
Cal / adjust
assemble
ND-fltr assy
qualified
ND motor
assy
…
Adjust filter to motor & cal
XWP motor
assy
4022.656.13051
4022.656.13081
4022.656.13091
4022.656.13221
4022.656.13121
4022.656.13141
Shaft.
VQI 3
VQI 3
VQI 3
VQI 3
VQI 3
VQI 2
Filter
3 of 4
Filter
4 of 4
Filter
2 of 4
Filter
1 of 4
Wheel frame
R
R
R
R
S
S
656. 13391
VQI 2
VQI 2
VQI 2
VQI 2
VQI 2
VQI 3
VQI 3
VQI 2
VQI 2
∑10k€
∑2.5k€
∑7k€
∑3.5k€
∑3.5k€
∑3.5k€
∑3k€
COG
LPPS-CL
VQI 3
VQI 3
∆-30k€
∑10k€
S
S
S
Diagn: ACPD Sens&Act
Diagn: TBD
Diagn: TBD
Diagn: TBD
….
cabling
Encoder
bearing
FAN
R
cabling
cabling
Encoder
NTC
PCA
S
H/V-Fiber
To shutter
S
D000467306
Qualify @
supplier
Qualify @
supplier
Interpolator
Interpolator
4022.666.47751
Fan Exhaust
S
holder
Bond optic to holder (6x)
LPPS2CL
Fiber assy
Frame
mirror1
mirror2
BS1 / BS2
Entrance
Fiber adj assy
Exit
Fiber adj
assy
Lens (6x)
NTC
Optical Assy
Adjust & Cal
BUY?
S
656. 13291
∑HW = 43k€

## Slide 10: YS SE CL SHORT/LONG WAVE WHEEL ASSY
Inertia driven design to meet timing spec’s
Notes: In order to switch fast we cut the filters in smaller pieces and glue them on a wheel. The challenge was balancing between low inertia and long filters. As low inertia has impact on timing and long filters on steepness of the filters, as there are then less nm in optical spot.

## Slide 11: Nominal wavelength positions on filter wheel
With geometrical tolerances
2mm
5mm
410nm
652.5nm
≅531nm
652.5nm
2mm
652.5nm
5mm
895nm
≅774nm
2mm
5mm
410nm
652.5nm
≅531nm
652.5nm
2mm
652.5nm
5mm
895nm
≅774nm
LWP
SWP

## Slide 12: YS SE CL XWP MOTOR ASSY
Motor layout
Notes: Switching from one wavelength to another < 45ms. (180°)
Motors placed next to each other and rotate in opposite directions to compensate disturbance torques. (0.25Nm)
Motor frame air-cooled
Closest position of the filters wheels is 0.5mm

## Slide 13: YS SE CL XWP MOTOR ASSY
First filter wheels mounted on motor assy

## Slide 14: YS SE CL OPTICS ASSY
Optical layout around moving filter wheels
Notes: 12 optical components, only one adjusted in XY. Mirrors and BS positioned on pads with springs. Lenses glued in lens cells and placed on locating pins.
All 4 fiber connectors adjusted as well.

## Slide 15: YS SE CL OPTICS ASSY
Fully assembled on main tool

## Slide 16: YS SE CL OPTICS ASSY
Lens cell with leaf springs for thermal shocks during transport

## Slide 17: YS SE CL OPTICS ASSY
Placing of first mirror

## Slide 18: YS SE CL OPTICS ASSY
Springs holding mirrors and beam splitters on pads

## Slide 19: YS SE CL ND FLTR-MTR ASSY
Attenuation of the light

## Slide 20: YS SE CL HOUSING ASSY

## Slide 21: YS SE CL MODULE
Fully assembled module

## Slide 22: Spare parts accessibility
All spare parts accessible from the front
Cooling fan
Attenuator (5 ND filters)
Motor block
Active PCA
Calibration source
Interpolators
Spectro meter
Internal fibers
Linear variable filters

## Slide 23: FUMO MTD and optics

## Slide 24: Assembled motor blocks

## Slide 25: Housing fittest and optical frame 3D measuring

## Slide 26
Separate submodules NX models

## Slide 31: 4022.666.35152 YS SE XWP MOTOR ASSY

## Slide 32: 4022.666.35201 YS SE OPTICS ASSY

## Slide 33: 4022.666.35451 YS SE CL ND FLTR-MTR ASSY

## Slide 34: 4022.666.35501 YS SE CL HOUSING ASSY

## Slide 35: 4022.656.13051 YS SE LONG WAVE WHEEL ASSY 4022.656.13061 YS SE SHORT WAVE WHEEL ASSY

## Slide 36: 4022.668.41861 YS CLM EXHAUST FAN ASSY

## Slide 37: 4022.668.41901 YS CLM INTERPOLATOR ASSY

## Slide 38: 4022.668.41831 YS CLM CALIBRATION ASSY

## Slide 39: 4022.668.41881 YS CLM CONVERTER ASSY

## Slide 40: 4022.666.35401 YS SE CL SPECTROMETER ASSY

## Slide 41: 4022.666.58971 YS SE CL FIBER IN ASSY

## Slide 43: Airflow cooling