## Slide 1
RLKM
25-06-2018
Sprint 106
FSS HorizontalVerticalOutputAdjustment

## Slide 2: Document history
Maintainer | Modified date | Status | Comment
JUDI | 25/06/2015 |  | Creation
RLKM | 24/12/2018 | For Review | Update

## Slide 3: Reference documents
Document ID | Title
D000350053/01-EDS EPDS | D000350053/01-EDS EPDS YS350E H/V Output Adjustment and Performance Tasks
D000586899/03-EDS | D000586899/03-EDS- EDS YS1375 task 402 - HV Output Adjustment task

## Slide 4: Open issues
ID | Description | AIR ID | Action Holder | Status
1 | 1375 Task 402 H V adjustment must be adapted for Pupil camera and to maximize H &V iso mimizing delta H&V | P340972 | RLKM | Closed
2 |  |  |  | 
3 |  |  |  | 
4 |  |  |  | 

## Slide 5: Contents
Scope
Goal & Context
Procedure
Input / Output
Spec, SPC and SDT parameters
Task dependencies and sequence

## Slide 6: Scope
The information in this document is applicable as of:
AIR ID | Release | Service Pack
P181201 | ys300c | ys300c_sp_04
P340972 | ys300c | ys300c_sp_19

## Slide 7: Goal & Context
The YS350 system has two Illumination fibers which each are equipped with a shutter-box assy. Apart from their shutter function (to select illumination polarization state by opening or closing the required shutter), these shutter-boxes have a function to compensate intensity differences between the two polarizations at Darkfield camera level.
Update in YS1375 : The used camera is changed into Pupil Camera instead of Darkfield camera.

## Slide 8: Task location in test manager(Update in SP19)
The task can be found by typing #402

## Slide 9: Procedure
Determine the optimal relative output difference for H/V polarized light
Acquire camera images for H/V polarization
  Determine average intensity levels on dark field image for negative and positive first order images
  Calculate relative difference between H/V polarized images.
  Store maximum reached intensity levels for both H/V polarized light
Spec is valid whenever relative difference levels between H/V are within -2 and +2 percent.
Note: camera type is different according to machine type
- YS1375 : Pupil camera
- YS375 : Dark field camera

## Slide 10
UI Update in SP19 :

## Slide 11: Input/Output
Input settings
Output report
New diagnostic settings

## Slide 12: Input settings
YS375 :
  Attenuator: Attenuator to be used for measurement ‘factor10’
  Repeat count: The number of repeating measurement sequences
  SensingLight: Wavelength to be used for measurement ‘550NB’
YS1375:

## Slide 13: Output report values
Integration time used
Final intensity level for H Polarized light
Final intensity level for V Polarized light
Maximum measured intensity level for H Polarized light
Maximum measured intensity level for V Polarized light
Relative intensity difference

## Slide 14: Output report
YS375 :

## Slide 15: Output report
YS1375 :

## Slide 16: New diagnostic settings
Optimal integration time
Number of H Polarized images used
Number of V Polarized images used
Relative intensity level, Latest intensity levels H/V, Maximum measured intensity level H/V of camera images
Relative energy sensor reading level, Latest energy sensor reading H/V, Maximum measured energy sensor reading H/V

## Slide 17: Spec, SPC and SDT parameters
New spec parameters available in SpecChecker
New SPC parameters available

## Slide 18: New spec parameters
RelativeIntensityDifference

## Slide 19: New SPC parameters
NA

## Slide 20: Task dependencies and sequence
Dependencies of task
  NA
Position of task in build sequence
  NA