## Slide 1
YANN/TINR
W1842
LF YieldStar Sustaining Source Project Execution Summary

## Slide 2: 2018 Q1 ~ Q3 Achievement (Source Project)
Task 616
DC offset
Task 624
DC offset
Task 618
Irradiance
Task 620
WL fine
Task 627
WL fine
Task 629
Spectrometer
Task 617
Non-linearity
Task 605
WL home offset
Task 602
Enc Pos to WL
Task 641
ND filter
Task 601
ND filter
Task 603
WL offset
Task 630
CLM spectral
Task 625
Enc Pos to WL
CL Module
External Spectrometer
New task VHV
new task
ACE
existing task
< SP18
Task 613
WL coarse
Task 610
Spectral
Task 600
WL coarse
Task 623
DC Offset
Task 604
DC offset
Task 611
Non-linearity
Task 612
Irradiance
Task 619
WL fine
Task 628
Spectrometer
Task 626
WL fine
Internal Spectrometer
Task 630 has 603 included
ACE contributes 16 out of 24 blocks for CLM SP18 delivery!
Fig. Full YS(1)375 CL Module Tasks Package
Notes: CLM source project contains 3 parts: internal/CLM/external….
Blocks with empty rectangle are existing tasks.
Others are newly introduced tasks.
Among them ACE participates……….
6 tasks for internal spectrometer/ 5 tasks for CL module/ 5 tasks for external spectrometer
Therefore, ACE contributes totally 16 tasks for CLM SP 18 delivery.

## Slide 3: 2018 Q1 ~ Q3 Achievement (Source Project)
CLM Q-Tool
  DC offset calibration (616)
  Non-linearity calibration (617)
  Irradiance calibration (618)
  Wavelength fine calibration (620/627)
Internal spectrometer
  DC offset calibration (604/623)
  Non-linearity calibration (611)
  Irradiance calibration (612)
  Wavelength fine calibration (619/626)
CLM calibrations
  ND filter calibration (601/641)
  Wavelength offset calibration (603)
  Encoder-wavelength fine calibration (625)
  Spectra performance calibration (630)
1
2
3
Support YS(1)375 CLM R3 (design verification) plan for SP18 delivery.
WK

## Slide 4: 2018 Q1 ~ Q3 Achievement (Source Project)
KTD
  CL module introduction
  DC offset calibration (604/623)
  Non-linearity calibration (611/617)
  Irradiance calibration (612/618)
  Wavelength fine calibration (619/626)
  ND filter calibration (604/641)
Shutter box
  Develop shutter open/close inline performance (123i)
  YS1375 HV intensity adjustment – P367213
  Task 123i failed on YS1375 SP17 – P373323
4
5
WK

## Slide 5: 2019 Ambition (Source Project)
Visit VHV for supporting CLM tasks final qualification and OIM knowledge transfer (WK1843 ~WK1844)
YS380 source NPI design engagement.
YS(1)375 Source ownership after KD14.
  CLM sustaining.
  OIM sustaining.
  LPPS sustaining.

## Slide 6
Q & A

## Slide 7: YS(1)375 Source
MOTOR FRAME
5x ND
1x slot
3x spare
OPTICAL BASE FRAME
LWP
SWP
LPPS
Mixing fiber
Z
Y
≅4%
≅ 4%
Spectrometer
H
V
Cal. source
F1
L1
M1
L2
L3
BS1
M3
L4
F3
F2
F4
L5
BS2
ND1
M2
L6
ND wheel
Spectrometer calibrations
Task#605:xWP home offset calibration
Task#602:xWP encoder position to wavelength calibration
Task#601: ND wheel position calibration
Task #123i
Shutter box