## Slide 1
Tina
wk1840
KTD CLM Spectrometer Tasks
#604 IlluminationSpectrometerDCOffsetCalibration
#623 IlluminationSpectrometerDCOffsetPerformance

## Slide 2: CLM Optical Structure
MOTOR FRAME
5x ND
1x slot
3x spare
OPTICAL BASE FRAME
LWP
SWP
LPPS
≅4%
≅ 4%
Internal Spectrometer
H
V
Cal. source
F1
L1
M1
L2
L3
BS1
M3
L4
F3
F2
F4
L5
BS2
ND1
M2
L6
ND wheel
Task#604: Spectrometer DC Offset calibration
Main path
Spectrometer path
Calibration path
Must operate task #605 in advance before starting task #604

## Slide 3: The DC offset origin
The dark current offset has DC and AC components:
Dark current would increase as the temperature increases
Electronic offset such as programmed DC offset
Integrated dark photodiode current.
Dark current shot noise
Read-noise
AC
DC
Position
Integration time
Temperature
Notes: AC
Shot noise is associated with the random arrival of photons which obeys a square root behavior w.r.t the dark-current),
read-noise comes from the pixel electronics such as in-pixel amplifier and read-out chain.

## Slide 4: Linear regression
The fitted equation
Minimize the error sum of squares (SSE)
Through some calculation, obtain the coefficients b0, b1,
Notes: In statistics, linear regression is a linear approach to modelling the relationship between a scalar response (or dependent variable) and one or more independent variables.

## Slide 5: Spectrometer DC Offset Calibration In Task #604
Dark current offset is an offset of the spectrum in dark conditions.
Task 604
DC offset
Task 600
WL coarse
Task 605
WL home offset
Task 623
DC Offset

## Slide 6: Spectrometer DC Offset Calibration In Task #604
Apply linear regression to obtain the coefficients, C and S, of the fitting line.
Subtract the fitted spectra from the measured spectra to get the corrected spectra I’
Calculate the slope scaling factor(using the last 12 dark pixels)
2.7
For a specific pixel
100 acquisitions for each integration time
Blueuncorrected
Redcorrected

## Slide 7: How to perform task #604
1. Equipment constant  Sensing  Spectrometer
Make True
2. Select task #604, check the default input setting and commit ECs if there is no problem
Default input setting | 
Minimal integration time[ms] | 0.03
Elbow integration time [ms] | 0.25
Maximal integration time [ms] | 2.5
Number of points [min, elbow] | 12
Number of point(elbow, max] | 23
Duration for calibration [s] | 120
Notes: The temperature fluctuation is of time scale of 120s.

## Slide 8: How to verify task #604
1. Select task #623, check input setting, run the task.
Input setting of task #623 | 
Minimal integration time[ms] | 0.03
Maximal integration time [ms] | 2.5
Number of points | 50
Duration for calibration [s] | 10
< 50.
2. Check the following specification:
Uniformity of pixels across each acquisition at given integration time
pixels

## Slide 9: Spectrometer DC Offset Calibration Validation
2.5 ms int. time | Before cal. | After cal.
Reports |  | 
Average Spectrum Histogram |  | 
 |  | Maximal integration time
Average DCOffset [DN16] | Min | 236.019
 | Mean | 597.893
 | Max | 678.794
Spatial Noise [DN16] |  | 169.838
 |  | Maximal integration time
Average DCOffset [DN16] | Min | -11.227
 | Mean | 1.603
 | Max | 13.624
Spatial Noise [DN16] |  | 9.598
~169
~9
Notes: Spatial noise is the 3 sigma of all pixels in a specific integration time.

## Slide 10: OCAP
Potential Problems | Possible Causes | Solutions
Pixel noise too high Frame-to-frame noise too high | Increased temperature Presence of stray light Spectrometer HW issue | Check CL module spectrometer temperature, ensure temp as close to 35̊C. Ensure spectrometer shielded properly from stray light during calibration. Check noise histogram to see if an increase in hot-pixels due to aging might have occurred. In such a case, the spectrometer might be defective.
SPEC: Maximal offset level noise (3∑) < 50
No OCAP for current version

## Slide 11
Q & A